Coating Adherence Verification,

Silq Surface Treatment

Quality Control Procedure

# Purpose

## The purpose of this document is to detail the quality control procedure used to verify the coating using the Silq Surface Treatment process remains adhered to the catheter. The coated catheter is rinsed with deionized water and the rinse is examined for the presence of coating using spectroscopy.

# Scope

## This procedure applies to catheters treated with the Silq Surface Treatment process.

# Documents

## Reference Documents

### QM.SLQ049	Workstation Practices SOP

### QC-E.SLQ015	Quality Control Qualification Procedure, Jenway 7205 Spectrophotometer

# Definitions

## Abbreviations / Acronyms

### cm: Centimeter

### DI: Deionized

### in: Inches

### mL: Milliliter

### nm: Nanometer

# Responsibilities

## Pathway MedTech (Pathway) is responsible for:

### Manufacturing, including final packaging and sterilization, of catheters in accordance with established manufacturing procedures

## Silq is responsible for:

### Performing crystal violet absorbance test in accordance with this procedure

### Documenting the results of the crystal violet absorption test

# Safety Information

## Gloves (non-latex) and safety eye wear are to be worn at all times during the execution of this procedure.

# Equipment / Materials / Supplies

## Spectrophotometer (E.SLQ015)

## UV Cuvette (S.SLQ011)

## DI Water

## Glass test tubes (S.SLQ010)

## Non-Latex Gloves

# Line Clearance and Setup

## Perform line clearance of the work area in accordance with Workstation Practices QM.SLQ049.

## Power up the spectrophotometer and allow to warm up for at least 15 minutes. Set the wavelength to 585nm.

## The spectrophotometer shall be qualified daily before use. Once the spectrophotometer is warmed up, qualify the spectrophotometer in accordance with QC-E.SLQ015.

# Procedure

## Secure a glass tube (S.SLQ010) and fill with DI water to the demarcated line.

## Take a dried, surface treated catheter and immerse it in the tube. Gently lift the catheter out of the tube and solution and then resubmerge it for a total of five times.

## Remove the catheter from the tube.

## Using a gloved finger to plug the top of the tube, invert the tube at least 3 times vigorously to ensure that any materials present in solution will be homogenously distributed.

## Pour out 2mL of the solution into a clean cuvette.

## Separately, pour 2mL of pure DI water into a UV cuvette (S.SLQ011) and use this cuvette to blank the spectrophotometer at 440 nm.

## Measure the absorbance value of the sample cuvette and document the results in Attachment 1.

# Acceptance Criteria

## The absorbance value shall be a less than or equal to 0.02.

# Attachments

## Attachment 1: Coating Adherence Verification Datasheet

# Attachment 1: Coating Adherence Verification Datasheet

## Tester (Print name & Sign): __________________________________________________________   Date of Test: ____________________

# General Product Information

| Part Number: |  | Part Description: |  |
| --- | --- | --- | --- |
| Part Revision: |  | Lot #: |  |

# Equipment Traceability

| Part Number | Description | Equipment ID | Last qualified on* |
| --- | --- | --- | --- |
| E.SLQ015 | Jenway 7205 Spectrophotometer |  |  |

* Attach QC Qualification Datasheet

## Summary of Results

| Catheter Sample # | Result | Acceptance criteria | Meets Criteria? |
| --- | --- | --- | --- |
| 1 |  | ≤ .02 | Yes No |
| 2 |  | ≤ .02 | Yes No |
| 3 |  | ≤ .02 | Yes No |
| 4 |  | ≤ .02 | Yes No |
| 5 |  | ≤ .02 | Yes No |
| 6 |  | ≤ .02 | Yes No |
| 7 |  | ≤ .02 | Yes No |
| 8 |  | ≤ .02 | Yes No |
| 9 |  | ≤ .02 | Yes No |
| 10 |  | ≤ .02 | Yes No |
| 11 |  | ≤ .02 | Yes No |
| 12 |  | ≤ .02 | Yes No |
| 13 |  | ≤ .02 | Yes No |
| 14 |  | ≤ .02 | Yes No |
| 15 |  | ≤ .02 | Yes No |
| 16 |  | ≤ .02 | Yes No |
| 17 |  | ≤ .02 | Yes No |
| 18 |  | ≤ .02 | Yes No |
| 19 |  | ≤ .02 | Yes No |
| 20 |  | ≤ .02 | Yes No |
| 21 |  | ≤ .02 | Yes No |
| 22 |  | ≤ .02 | Yes No |

## OVERALL TEST RESULTS (circle one):  PASS  /  FAIL

Reviewed By (Sign and Date):_____________________________________________________________