# Purpose

## The purpose of this document is to identify the associated risk with the intended use of the SILQ ClearTract Foley Catheter.

# Scope

## The product numbers associated with the SILQ ClearTract Foley Catheter is indicated in the following table:

| Table 1: SILQ ClearTract Catheter Products |  |
| --- | --- |
| ## Part No. | ## Description |
| ## HDX-1410 | ## 2-way 100% Silicone ClearTract Catheter, 14 Fr/10ml |
| ## HDX-1605 | ## 2-way 100% Silicone ClearTract Catheter, 16 Fr/5ml |
| ## HDX-1810 | ## 2-way 100% Silicone ClearTract Catheter, 18 Fr/10ml |
| ## HDX-160510 | ## 2-way 100% Silicone ClearTract Catheter, 16 Fr/5-10ml |
| ## SLQ-1610 | ## 2-Way 100% Silicone ClearTract Catheter, 16Fr/10ml |
| ## SLQ-211410 | ## 2-way 100% Silicone ClearTract Catheter, 14 Fr/10ml |
| ## SLQ-211605 | ## 2-way 100% Silicone ClearTract Catheter, 16 Fr/5ml |
| ## SLQ-211610 | ## 2-way 100% Silicone ClearTract Catheter, 16 Fr/10ml |
| ## SLQ-211810 | ## 2-way 100% Silicone ClearTract Catheter, 18 Fr/10ml |
| ## SLQ-211410 | ## 2-way 100% Silicone ClearTract SPT Catheter, 14 Fr/10ml |
| ## SLQ-211605 | ## 2-way 100% Silicone ClearTract SPT Catheter, 16 Fr/5ml |
| ## SLQ-211610 | ## 2-way 100% Silicone ClearTract SPT Catheter, 16 Fr/10ml |
| ## SLQ-211810 | ## 2-way 100% Silicone ClearTract SPT Catheter, 18 Fr/10ml |

## The intended use of this device is the following:

### The 2-Way 100% Silicone ClearTract Catheter is intended for drainage of the urinary tract. Catheterization is generally accomplished by inserting the catheter through the urethra and into the bladder. However, drainage is sometimes accomplished by suprapubic placement or other placement of the catheter, such as nephrostomy tract. Intended population is adults and pediatrics.

# References

## QP-004 – Risk Management

## RM-0018 – Risk Management Plan, SILQ ClearTract Foley Catheter

# Risk Tables Definitions

## # – Enumerate each Hazard Category with a sequential number

## Hazard Category – Describes the general grouping of a hazard(s) / hazardous situation(s) that is being evaluated for risk

## PH# – Enumerate each Potential Hazard with a sequential number, using the format {Hazard Category #}.{PH#}, (e.g. 1.1, 1.2, 1.3)

## Potential Hazard (PH) – Source of harm

## Potential Harm – Injury or damage to the health of people, or damage to property or the environment

## Severity (S) – Measure of the possible consequences of a hazard.

## Potential Failure Mode (FM) – The potential consequences of the failure

## Risk Controls - Process through which decisions are reached and protective measures are implemented for reducing risk to, or maintaining risk within specified levels.

### Risk Controls1 – Initial risk controls that are incorporated.

### Risk Controls2 – Secondary risk controls, also known as “Mitigation”.

## Verification – Evidence source of linkage that initial (i.e., Risk Controls1) risk controls have implemented, verified, or validated.

## Probability – Likelihood of the hazard occurring

### Probability1 (P1) – The probability assessment is post the implementation/ verification of initial risk controls (i.e., Risk Controls1)

### Probability2 (P2) – The probability assessment is post the implementation/ verification of the secondary risk controls (i.e., Risk Controls2)

## Risk Rating – Residual risk determination based on the concept of ‘As Far As Possible (AFAP)’ approach. The AFAP approach provides three (3) possible outcomes, which are 1) the residual risk exceeds the risk acceptability (RED), 2) the residual risk is acceptable to the risk acceptability criteria (GREEN), and 3) the residual risk is between the two (2) states indicated and additional controls are suggested to lower the risk to the lowest practical level, without economically considerations, based on management decision (YELLOW).

### Risk Rating1 (R1) – This calculation is post the implementation of initial risk controls (i.e., Risk Controls1)

### Risk Rating2 (R2) – This calculation is post the implementation of the secondary risk controls (i.e., Risk Controls2)

# Identify Characteristics that could Impact on Safety

| Table 2: : Questions from ISO-14971:2012 Annex C w/ Answers |  |  |
| --- | --- | --- |
| # | ISO-14971:2012 Annex C Question | Answer |
| 1 | What is the intended use and how is the medical device to be used? | Refer to 2.2 |
| 2 | Is the medical device intended to be implanted? | No |
| 3 | Is the medical device intended to be in contact with the patient or other persons? | Yes |
| 4 | What materials or components are utilized in the medical device or are used with, or are in contact with, the medical device? | Silicone |
| 5 | Is energy delivered to or extracted from the patient? | No |
| 6 | Are substances delivered to or extracted from the patient? | Yes, urine |
| 7 | Are biological materials processed by the medical device for subsequent re-use, transfusion or transplantation? | No |
| 8 | Is the medical device supplied sterile or intended to be sterilized by the user, or are other microbiological controls applicable? | Yes, supplied sterile by EO. |
| 9 | Is the medical device intended to be routinely cleaned and disinfected by the user? | No |
| 10 | Is the medical device intended to modify the patient environment? | No |
| 11 | Are measurements taken? | Yes, measurements of urine. |
| 12 | Is the medical device interpretative? | No |
| 13 | Is the medical device intended for use in conjunction with other medical devices, medicines or other medical technologies? | The device has been designed to be used with a urine collection bag. |
| 14 | Are there unwanted outputs of energy or substances? | No |
| 15 | Is the medical device susceptible to environmental influences? | No |
| 16 | Does the medical device influence the environment? | No |
| 17 | Are there essential consumables or accessories associated with the medical device? | No |
| 18 | Is maintenance or calibration necessary? | No |
| 19 | Does the medical device contain software? | No |
| 20 | Does the medical device have a restricted shelf-life? | Yes, refer to #23 |
| 21 | Are there any delayed or long-term use effects? | No |
| 22 | To what mechanical forces will the medical device be subjected? | The device is subjected to external tensile forces imposed by the urine content in the collection bag, Internal pressure exerted by urine within the lumen, external compressive forces of the bladder and urethra, internal forces from the filled balloon. |
| 23 | What determines the lifetime of the medical device? | Product and sterile barrier performance |
| 24 | Is the medical device intended for single use? | Yes |
| 25 | Is safe decommissioning or disposal of the medical device necessary? | Yes |
| 26 | Does installation or use of the medical device require special training or special skills? | No |
| 27 | How will information for safe use be provided? | Instructions for use |
| 28 | Will new manufacturing processes need to be established or introduced? | No |
| 29 | Is successful application of the medical device critically dependent on human factors such as the user interface? | No |
| 29.1 | Can the user interface design features contribute to use error? | No |
| 29.2 | Is the medical device used in an environment where distractions can cause use error? | No |
| 29.3 | Does the medical device have connecting parts or accessories? | Yes, urine collection bag, Luer-lock syringe |
| 29.4 | Does the medical device have a control interface? | No |
| 29.5 | Does the medical device display information? | No |
| 29.6 | Is the medical device controlled by a menu? | No |
| 29.7 | Will the medical device be used by persons with special needs? | No |
| 29.8 | Can the user interface be used to initiate user actions? | No |
| 30 | Does the medical device use an alarm system? | No |
| 31 | In what way(s) might the medical device be deliberately misused? | None |
| 32 | Does the medical device hold data critical to patient care? | No |
| 33 | Is the medical device intended to be mobile or portable? | Yes |
| 34 | Does the use of the medical device depend on essential performance? | No |

# Risk Acceptability Determination

## The criteria of risk acceptability will be determined with a Risk Priority Number (RPN) and on the concept of ‘As Far As Possible’ (AFAP) approach.  Refer to the Risk Rating definition in §4.0.

## There are two (2) components used in determining risk acceptability, which are Severity and Probability. Refer to the following tables:

| Table 3: Severity Ratings |  |
| --- | --- |
| Index | Description |
| 1 | No patient effect |
| 2 | Minor injury that resolves without treatment |
| 3 | Moderate injury that may require treatment |
| 4 | Serious injury requiring medical intervention |
| 5 | Life threatening injury, disability, death |

| Table 4: Probability Ratings |  |
| --- | --- |
| Index | Description |
| 1 | Improbable - <0.001% |
| 2 | Remote – 0.001% to .01% |
| 3 | Occasional – 0.01% to 0.1% |
| 4 | Likely – 0.1% to 1% |
| 5 | Frequent - >1% |

## Risk acceptability will be determined per the following risk index:

| Table 5: Risk Rating Table |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Severity | 5 | 5 | 10 | 15 | 20 | 25 |
|  | 4 | 4 | 8 | 12 | 16 | 20 |
|  | 3 | 3 | 6 | 9 | 12 | 15 |
|  | 2 | 2 | 4 | 6 | 8 | 10 |
|  | 1 | 1 | 2 | 3 | 4 | 5 |
|  |  | 1 | 2 | 3 | 4 | 5 |
|  |  | Probability |  |  |  |  |

# Risk Assessment Tables

| Table 6: SILQ ClearTract Foley Catheter |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| # | Hazard Category | PH# | Potential Hazard | Potential Harm | S | Potential Failure Mode | Risk Controls1 | Verification | P1 | R1 | Risk Controls2 | P2 | R2 |
| 1 | Acoustic Energy | The medical device does not have any acoustic energy hazards. |  |  |  |  |  |  |  |  |  |  |  |
| 2 | Electromagnetic Energy | The medical device does not have any electromagnetic energy hazards. |  |  |  |  |  |  |  |  |  |  |  |
| 3 | Mechanical Energy | The medical device does not have any mechanical energy hazards. |  |  |  |  |  |  |  |  |  |  |  |
| 4 | Potential (stored) Energy | The medical device does not have any potential energy hazards. |  |  |  |  |  |  |  |  |  |  |  |
| 5 | Radiation Energy | The medical device does not have any radiation energy hazards. |  |  |  |  |  |  |  |  |  |  |  |
| 6 | Thermal Energy | The medical device does not have any thermal energy hazards. |  |  |  |  |  |  |  |  |  |  |  |
| 7 | Biological | 7.1 | Biological Contamination | Fever/ infection to patient | 3 | Device is not provided sterile/ insufficient sterilization process | Device will be terminally sterilized with validated method (Design) | Sterilization validation TR-0075, TR-HDX-0009 | 1 | 3 | N/A | N/A | N/A |
|  |  |  |  |  |  | Device is reused | Warnings statement for single-use (Information) | IFU SLQ-4006 | 2 | 3 | N/A | N/A | N/A |
|  |  |  |  |  |  | Inadequate sterile barrier system | Sterile barrier system will be designed for material that are commonly used for single-use medical device (Design) | Packaging Validation TR-0074 | 1 | 3 | N/A | N/A | N/A |
|  |  |  |  |  |  |  |  | Stability TR-0073 |  |  |  |  |  |
|  |  |  |  |  |  |  | Warning will be incorporated for inspecting packaging prior to use as well not using the device if packaging is damaged (Information) | IFU SLQ-4006 | 2 | 6 | N/A | N/A | N/A |
|  |  |  |  |  |  | Use beyond expiration date | Warning statement (Information) | IFU SLQ-4006 | 2 | 6 | N/A | N/A | N/A |
|  |  | 7.2 | Biological Contamination | Fever/ infection to clinician | 3 | Proper handling and disposal are not conducted during clinical application | Warning statement (Information) | IFU SLQ-4006 | 2 | 6 | N/A | N/A | N/A |
| 8 | Chemical | 8.1 | Exposure to known allergens | Allergic reaction | 3 | Materials (e.g., latex, rubber) that is known to cause allergic response | Material will be selected that do not cause allergic response (Design) / Biocompatibility Testing | Design verification – biocompatibility TR-0071 | 1 | 3 | N/A | N/A | N/A |
|  |  | 8.2 | Exposure to toxic materials / False positives in the blood work evaluation / red blood cell damage | Toxicity | 3 | Materials leach during clinical application | Materials will be selected that are non-toxic (Design) | Biocompatibility Test Report TR-0071 | 1 | 3 | N/A | N/A | N/A |
|  |  |  |  |  |  | Excessive ethylene oxide residuals from sterilization process | Device will be released for distribution with acceptable residual levels (Design) | Sterilization Validation TR-0075, TR-HDX-0009 |  |  |  |  |  |
|  |  | 8.3 | Procedure complication – device inoperative | Delay in treatment | 3 | Material degrades over usable life | Material selection (Design) | Stability evaluation TR-0073 | 1 | 3 | N/A | N/A | N/A |
|  |  |  |  |  |  |  | Sterilization method selection (Design) |  |  |  |  |  |  |
| 9 | Immunological | 9.1 | Exposure to non-biocompatible materials | Cell Damage | 3 | Patient contact materials are not biocompatibility | Material selection (Design) | Biocompatibility Test Report TR-0071 | 1 | 3 | N/A | N/A | N/A |
|  |  | 9.2 | Exposure to bacterial endotoxins | Fever / pyrogenic response | 3 | Contacting material causes pyrogenic response | Material selection (Design) | Design verification – biocompatibility (Material mediated pyrogenicity) TR-0071 | 1 | 3 | BET testing / lot WI-0021 | 1 | 3 |
| 10 | Function (Operational) | 10.1 | Procedure complication – device inoperative | Delay in treatment | 2 | Balloon Burst – Material or dimensional issues (wall thickness) | Material selection (Design) Dimensional Assessment (Design) | Design verification – Balloon Integrity (Resistance to Rupture) TR-0072, TR-HDX-0001 | 1 | 2 | N/A | N/A | N/A |
|  |  |  |  |  | 2 | Balloon Burst – Material or dimensional issues (wall thickness) | Material selection (Design) Dimensional Assessment (Design) | Design verification – Balloon Integrity (Resistance to Rupture) TR-0072, TR-HDX-0001 | 1 | 2 | N/A | N/A | N/A |
|  |  |  |  |  | 1 | Balloon deflates or does not inflates – Valve malfunctions or leaks (wall thickness issue) | Valve selection (Design) Dimensional Assessment (Design) | Design verification – Balloon Volume Maintenance TR-0072, TR-HDX-0001 | 1 | 1 | N/A | N/A | N/A |
|  |  |  |  |  | 1 | Balloon does not inflate – Valve malfunctions | Valve selection (Design) | Design verification TR-0072, TR-HDX-0001 | 1 | 1 | N/A | N/A | N/A |
|  |  |  |  |  | 2 | Catheter breaks under normal tensile stresses – Material or dimensional issues | Valve selection (Design) Dimensional Assessment (Design) | Design verification – Inflated Balloon Response to Traction TR-0072, TR-HDX-0001 | 1 | 2 | N/A | N/A | N/A |
|  |  |  |  |  | 1 | Connection issues – Syringe (inflation, irrigation); Collection bag (drainage) | Connection will be standardized to what is incorporated with marketed catheters (Design) | Design verification TR-0072, TR-HDX-0001 | 1 | 1 | N/A | N/A | N/A |
|  |  | 10.2 | Procedure complication – downgraded performance | Delay in treatment | 2 | Poor drainage – Insufficient holes and lumen diameter, Lumen collapses (Materials or dimensional issues) | Material selection (Design) Dimensional Assessment (Design) | Design verification – Flow Rate though Drainage Lumen TR-0072, TR-HDX-0001 | 1 | 2 | N/A | N/A | N/A |
|  |  |  |  |  | 2 | Poor irrigation - Insufficient holes and lumen diameter, Lumen collapses (Materials or dimensional issues) | Material selection (Design) Dimensional Assessment (Design) | Design verification TR-0072, TR-HDX-0001 | 1 | 2 | N/A | N/A | N/A |
|  |  | 10.3 | Blunt / rough surface | Urinary tract damage | 3 | Tip geometry – Damage during insertion | Tip design will be standardized to what is incorporated with marketed catheters (Design) | First Article TR-0072, TR-HDX-0001 | 1 | 3 | N/A | N/A | N/A |
|  |  |  |  |  | 3 | Irrigation pressure too high – Insufficient hole geometry | Valve selection (Design) Dimensional Assessment (Design) | Design verification TR-0072, TR-HDX-0001 | 1 | 3 | N/A | N/A | N/A |
|  |  | 10.4 | Procedure complication – Difficult to remove | Extended procedure / Urinary tract damage | 3 | Balloon will not deflate – Valve malfunctions | Valve selection (Design) | Design Verification - Deflation Reliability (Failure to Deflate) TR-0072, TR-HDX-0001 | 1 | 3 | N/A | N/A | N/A |
| 11 | User Error (Operational) | 11.1 | Procedure complication – device inoperative | Delay in treatment | 1 | Catheter dislodges for patient – insufficient balloon inflation | Warning and limitations statements (Information) Define balloon fill volume | IFU SLQ-4006 | 2 | 2 | N/A | N/A | N/A |
|  |  | 11.2 | Partial / leaking connection to drain bag | Rash/Infection – Skin | 2 | Inadequate connection with collection bag | Product instruction for connection | IFU SLQ-4006 | 2 | 4 | N/A | N/A | N/A |
|  |  | 11.3 | Unskilled/ untrained clinician | Patient discomfort | 2 | Clinicians – does not read some or all information in IFU | Product labeling will clearly state to read the IFU prior to performing clinical application (Information) | Product Labeling SLQ-4000, SLQ-4001, SLQ-4006 | 2 | 4 | N/A | N/A | N/A |
|  |  |  |  |  |  |  | Warning and limitations statements (Information) | IFU SLQ-4006 |  |  |  |  |  |
|  |  |  |  |  | 3 | Balloon not deflated or partially deflated during removal | Warning and limitations statements (Information) | IFU SLQ-4006 | 2 | 6 | N/A | N/A | N/A |
|  |  |  |  |  | 3 | Excessive tensile – Catheter not secure, weight of collection bag pulls on catheter | Warning and limitations statements (Information) | IFU SLQ-4006 | 2 | 6 | N/A | N/A | N/A |
|  |  |  |  |  | 2 | Balloon burst - excessive balloon inflation | Warning and limitations statements (Information) Define balloon fill volume | IFU SLQ-4006 | 2 | 4 | N/A | N/A | N/A |
|  |  |  |  |  | 2 | Excessive irrigation pressure | Warning and limitations statements (Information) | IFU SLQ-4006 | 2 | 4 | N/A | N/A | N/A |
|  |  | 11.4 | Reuse of single use device | Fever/infection to patient | 3 | Device was reused | Warning and limitations statements (Information) | IFU SLQ-4006 | 2 | 6 | N/A | N/A | N/A |
|  |  | 11.5 | Use of a compromised device | Fever/infection to patient | 3 | Sterile barrier system was not inspected for damage | Warning and limitations statements (Information) | IFU SLQ-4006 | 2 | 6 | N/A | N/A | N/A |
|  |  | 11.6 | Use of expired device | Fever/infection to patient | 3 | Device used after expiration date or outside biocompatibility duration limits | Warning and limitations statements (Information) | IFU SLQ-4006 | 2 | 6 | N/A | N/A | N/A |
|  |  | 11.7 | Mishandled during disposal | Fever/infection to clinician | 3 | Occurs during disposal process | Warning and limitations statements (Information) | IFU SLQ-4006 | 2 | 6 | N/A | N/A | N/A |
| 12 | Labeling (Product Labeling) | 12.1 | Loss of traceability | Cannot identify the device or trace back to MFG | 1 | Incomplete or ambiguous instructions for use | Label and labeling will be designed to requirements indicated in regulation for the markets intended to be distributed (Information) | Label content review FILE-004318, FILE-008308 | 1 | 1 | N/A | N/A | N/A |
|  |  | 12.2 | Use of expired device / reuse of single use device | Fever/infection to patient | 3 | Incomplete or ambiguous instructions for use | Label and labeling will be designed to requirements indicated in regulation for the markets intended to be distributed (Information) | Label content review FILE-004318, FILE-008308 | 1 | 3 | N/A | N/A | N/A |
| 13 | Operating Instructions (Product Labeling) | 13.1 | Procedural Complication / misuse | Patient discomfort | 3 | Incomplete or ambiguous instructions for use | IFU will be designed to requirements indicated in regulation for the markets intended to be distributed (Information) | Label content review FILE-004318, FILE-008308 | 1 | 3 | N/A | N/A | N/A |
| 14 | Environmental | 14.1 | Deteriorated device / Sterile barrier failure | Patient health is compromised | 3 | Stored outside the prescribe environmental parameters | Warning statement (Information) | IFU SLQ-4006 | 2 | 6 | N/A | N/A | N/A |
|  |  |  |  |  |  | Packaging is not adequate to protect the device through transportation | Packaging will be designed to protect the device and its sterile barrier through transportation stresses (Design) | Design Verification (Distribution Evaluation) TR-0074 | 1 | 3 | N/A | N/A | N/A |

# Risk/ Benefit Assessment

## The following table indicates the risk/benefit analysis for the residual risks that were greater than the risk acceptability limit after all applicable controls/mitigations were applied:

| Table 7: Risk/Benefit Assessment for Residual Risk above Acceptability Limit for Process Hazard Assessment |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Table No. | Hazard Category | PH# | Potential Hazard / Harm | Potential Failure Mode | S | P2 | Risk/Benefit Analysis | Does the benefit outweigh the residual risk? |
| N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |

## There was no residual risk identified that was above the risk acceptability limit per §6.3.

# Conclusion

## All Identified risks regardless of acceptability were reduced as far as possible (AFAP). All risks, including residual risks and regardless of their dimension, were reduced “as far as possible” (AFAP). Risks have been reduced as low as possible by the following methods:

## Design

## Appropriate testing to ensure design criteria were met

## Protective measures (critical-to-quality)

## Manufacturing and quality control measures

## Based on a review of the risk-management process for the SILQ Foley Catheter described herein, the risk management plan based on ISO 14971:2019 has been implemented and the individual, cumulative and overall residual risk have been reduced as low as possible regardless of the magnitude and the benefit of using this device will outweigh risks posed by the device during usage.

# Change History

| ## Revision | ## Change Description |
| --- | --- |
| ## A | ## Initial release |
| ## B | ## Add HDX-160510, general updates, add verification references |
| ## C | ## Add SLQ-1610, add column Potential Harm to Table 6, general updates |
| ## D | ## Update Scope Table 1 and §2.2, add reference to RM Plan §3.2, Move 8.3 to 9.2, reclassify pyrogenic response under immunological risk |