# Purpose

## The purpose of this document is to identify the associated risk with the design of the components and their interactions with each other.

# Scope

## The product numbers associated with the SILQ ClearTract Foley Catheter is indicated in the following table:

| Table 1: SILQ ClearTract Catheter Products |  |
| --- | --- |
| ## Part No. | ## Description |
| ## HDX-1410 | ## 2-way 100% Silicone ClearTract Catheter, 14 Fr/10ml |
| ## HDX-1605 | ## 2-way 100% Silicone ClearTract Catheter, 16 Fr/5ml |
| ## HDX-1810 | ## 2-way 100% Silicone ClearTract Catheter, 18 Fr/10ml |
| ## HDX-160510 | ## 2-way 100% Silicone ClearTract Catheter, 16 Fr/5-10ml |
| ## SLQ-1610 | ## 2-Way 100% Silicone ClearTract Catheter, 16Fr/10ml |
| ## SLQ-211410 | ## 2-way 100% Silicone ClearTract Catheter, 14 Fr/10ml |
| ## SLQ-211605 | ## 2-way 100% Silicone ClearTract Catheter, 16 Fr/5ml |
| ## SLQ-211610 | ## 2-way 100% Silicone ClearTract Catheter, 16 Fr/10ml |
| ## SLQ-211810 | ## 2-way 100% Silicone ClearTract Catheter, 18 Fr/10ml |
| ## SLQ-211410 | ## 2-way 100% Silicone ClearTract SPT Catheter, 14 Fr/10ml |
| ## SLQ-211605 | ## 2-way 100% Silicone ClearTract SPT Catheter, 16 Fr/5ml |
| ## SLQ-211610 | ## 2-way 100% Silicone ClearTract SPT Catheter, 16 Fr/10ml |
| ## SLQ-211810 | ## 2-way 100% Silicone ClearTract SPT Catheter, 18 Fr/10ml |

# References

## QP-004 – Risk Management

## RM-0018 – Risk Management Plan, SILQ ClearTract Foley Catheter

# Risk Tables Definitions (ordered left to right)

## # – Enumerate each Description with a sequential number

## Description – Describes the operation that is being evaluated for risk

## Potential Failure Mode (FM) – The potential consequences of the failure

## Severity (S) – Measure of the possible consequences of a hazard.

## Potential Effect of Failure – All conceivable effects of the potential failure mode

## Potential Cause of Failure – All conceivable causes of the potential failure mode

## Risk Controls - Process through which decisions are reached and protective measures are implemented for reducing risk to, or maintaining risk within specified levels.

### Risk Controls1 – Initial risk controls that are incorporated.

### Risk Controls2 – Secondary risk controls, also known as “Mitigation”.

## Verification – Evidence source of linkage that initial (i.e., Risk Controls1) risk controls have implemented, verified, or validated.

## Probability – Likelihood of the hazard occurring

### Probability1 (P1) – The probability assessment is post the implementation/ verification of initial risk controls (i.e., Risk Controls1)

### Probability2 (P2) – The probability assessment is post the implementation/ verification of the secondary risk controls (i.e., Risk Controls2)

## Detection - The ability to screen out a hazard through verification activities (e.g., inspection).

### Detection1 (D1) – The detection assessment is post the implementation/ verification of initial risk controls (i.e., Risk Controls1)

### Detection2 (D2) – The detection assessment is post the implementation/ verification of the secondary risk controls (i.e., Risk Controls2)

## Frequency – A ‘Frequency Score’ will be calculated using the ratings from probability and detection components. The frequency equation is the following:   Note: The value calculated should be rounded up.

### Frequency1 (F1) – The frequency score calculated using Risk Controls1 and Detection1

### Frequency2 (F2) – The frequency score calculated using Risk Controls2 and Detection2

## Risk Rating – Residual risk determination based on the concept of ‘As Far As Possible’ (AFAP) approach. The AFAP approach provides three (3) possible outcomes, which are 1) the residual risk exceeds the risk acceptability (RED), 2) the residual risk is acceptable to the risk acceptability criteria (GREEN), and 3) the residual risk is between the two (2) states indicated and it can be an option to lower the risk to the lowest practical level without economic considerations (YELLOW).

### Risk Rating1 (R1) – This calculation is post the implementation of initial risk controls (i.e., Risk Controls1)

### Risk Rating2 (R2) – This calculation is post the implementation of the secondary risk controls (i.e., Risk Controls2)

# Risk Acceptability Determination

## The criteria of risk acceptability will be determined with a Risk Priority Number (RPN) and on the concept of ‘As Far As Possible (AFAP)’ approach. Refer to the Risk Rating definition in §4.0.

## There are three (3) components used in determining risk acceptability, which are Severity, Probability, and Detection. Refer to the following tables:

| Table 2: Severity Ratings |  |
| --- | --- |
| Index | Description |
| 1 | No patient effect |
| 2 | Minor injury that resolves without treatment |
| 3 | Moderate injury that may require treatment |
| 4 | Serious injury requiring medical intervention |
| 5 | Life threatening injury, disability, death |

| Table 3: Probability Ratings |  |
| --- | --- |
| Index | Description |
| 1 | Improbable - <0.001% |
| 2 | Remote – 0.001% to .01% |
| 3 | Occasional – 0.01% to 0.1% |
| 4 | Likely – 0.1% to 1% |
| 5 | Frequent - >1% |

| Table 4: Detection Ratings |  |
| --- | --- |
| Index | Description |
| 1 | Very High – Current controls are almost certain to detect the failure mode. |
| 2 | High – High likelihood that current controls will detect the failure mode. |
| 3 | Moderate – Moderate likelihood that current controls will detect the failure mode. |
| 4 | Low – Low likelihood that current controls will detect the failure mode. |
| 5 | Remote – Remote likelihood that current controls will detect the failure mode. |

## Risk acceptability will be determined per the following risk index:

| Table 5: Risk Rating Table |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Severity | 5 | 5 | 10 | 15 | 20 | 25 |
|  | 4 | 4 | 8 | 12 | 16 | 20 |
|  | 3 | 3 | 6 | 9 | 12 | 15 |
|  | 2 | 2 | 4 | 6 | 8 | 10 |
|  | 1 | 1 | 2 | 3 | 4 | 5 |
|  |  | 1 | 2 | 3 | 4 | 5 |
|  |  | =Square Root (Probability X Detection) |  |  |  |  |

Note: The purpose of using the square root of the probability and detection scores is to deemphasize these two (2) components, which places more emphasize on the severity component.

# Risk Assessment Tables

| Table 6: Product Risk Assessment |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| # | Component Description - Function | Potential Failure Mode | S | Potential Effects of Failure | Potential Causes of Failure | Risk Controls1 | Verification | P1 | D1 | F1 | R1 | Risk Controls2 | P2 | D2 | F2 | R2 |
| 1 | Catheter tube transfers urine from patient into drainage bag | Poor drainage / occluded lumen | 2 | Patient discomfort – device performance | Insufficient holes, lumen diameter, wall thickness | Design follows standard ASTM F623 | TR-0072, TR-HDX-0001 (flow test) | 1 | 2 | 2 | 4 | N/A | N/A | N/A | N/A | N/A |
|  |  | Catheter breaks | 2 | Procedural Complication – device failure | Wall thickness | Design follows standard ASTM F623 | TR-0072, TR-HDX-0001 (inflated balloon response to traction) | 1 | 2 | 2 | 4 | N/A | N/A | N/A | N/A | N/A |
|  |  |  |  |  | Performance degrades over time | Material Selection (Design) | Stability Evaluation TR-0073 |  |  |  |  |  |  |  |  |  |
|  |  | Oversized tube | 3 | Procedural Complication – Urethral damage | Wall thickness | Design follows standard ISO 20969 | TR-0072, TR-HDX-0001 (Balloon size and shaft size) | 1 | 2 | 2 | 6 | N/A | N/A | N/A | N/A | N/A |
| 2 | Catheter tube allows for irrigation | Poor irrigation/ occluded lumen | 1 | Procedural Complication – device failure | Insufficient holes, lumen diameter, wall thickness | Design follows standard ASTM F623 | TR-0072, TR-HDX-0001 (flow test) | 1 | 2 | 2 | 4 | N/A | N/A | N/A | N/A | N/A |
| 3 | Balloon inflates to secure catheter in place | Balloon leaks / deflates | 2 | Procedural Complication – Catheter dislodges | Wall thickness | Design follows standard ASTM F623 | TR-0072, TR-HDX-0001 (Balloon Volume Maintenance) | 1 | 2 | 2 | 4 | N/A | N/A | N/A | N/A | N/A |
|  |  |  |  |  | Performance degrades over time | Material Selection (Design) | Stability Evaluation TR-0073 |  |  |  |  |  |  |  |  |  |
|  |  | Balloon won’t inflate | 1 | Procedural Complication – No patient harm | Wall thickness | Design follows standard ASTM F623 | TR-0072, TR-HDX-0001 (Balloon Volume Maintenance) | 1 | 2 | 2 | 2 | N/A | N/A | N/A | N/A | N/A |
|  |  | Balloon bursts | 2 | Procedural Complication | Wall thickness | Design follows standard ASTM F623, Fill volume | TR-0072, TR-HDX-0001 (Balloon integrity – resistance to rupture) | 1 | 2 | 2 | 4 | N/A | N/A | N/A | N/A | N/A |
|  |  |  |  |  | Performance degrades over time | Material Selection (Design) | Stability Evaluation TR-0073 |  |  |  |  |  |  |  |  |  |
| 4 | Valve allows inflation and deflation of balloon and connects with a syringe | Leaking Valve – fail to seal | 2 | Procedural Complication - Catheter dislodges | Valve malfunction | Design follows standard ASTM F623 | TR-0072, TR-HDX-0001 (Balloon Volume Maintenance) | 1 | 2 | 2 | 4 | N/A | N/A | N/A | N/A | N/A |
|  |  |  |  |  | Performance degrades over time | Material Selection (Design) | Stability Evaluation TR-0073 |  |  |  |  |  |  |  |  |  |
|  |  | Balloon won’t deflate – valve fails to open | 3 | Procedural Complication – difficult to remove | Valve malfunction | Design follows standard ASTM F623 | TR-0072, TR-HDX-0001 (Deflation Reliability – failure to deflate) | 1 | 2 | 2 | 4 | N/A | N/A | N/A | N/A | N/A |
|  |  |  |  |  | Performance degrades over time | Material Selection (Design) | Stability Evaluation TR-0073 |  |  |  |  |  |  |  |  |  |
|  |  | Syringe won’t connect to valve | 1 | Complications in manufacturing assembly process | Not designed to Luer standards | Design follows standard ASTM F623 | TR-0072, TR-HDX-0001 (Verification Non-Interconnectable Connector) | 1 | 1 | 1 | 1 | N/A | N/A | N/A | N/A | N/A |
|  |  | Valve disconnects from catheter | 1 | Procedural Complication – device failure | Inadequate design | Design follows standard ASTM F623 | TR-0072, TR-HDX-0001 (Verification Non-Interconnectable Connector) | 1 | 2 | 2 | 2 | N/A | N/A | N/A | N/A | N/A |
| 5 | Pouch – Provide sterile barrier | Breach in sterile barrier | 3 | Non-sterile device | Inadequate seal | Material selection (Device) | Design verification (seal integrity) TR-0074 | 1 | 1 | 1 | 3 | N/A | N/A | N/A | N/A | N/A |
|  |  |  |  |  | Inadequate material (e.g., sterilization process, product weight, sharp edges) |  | Design verification (distribution evaluation) TR-0074 |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  | Performance degrades over time |  | Stability Evaluation TR-0073 |  |  |  |  |  |  |  |  |  |
| 6 | Information for Use (DFU) – Provides device information | Complication with device use/ intended purpose | 3 | Patient harm | Incomplete or ambiguous instructions for use | IFU will be designed to requirements indicated in regulation for the markets intended to be distributed (Information) | Label content review FILE-004318, FILE-008308 | 1 | 1 | 1 | 3 | N/A | N/A | N/A | N/A | N/A |
| 7 | Inner Carton – Protects and orients the pouched devices | Packaging is damaged | 3 | Device/ sterile barrier damage | Inadequate design | Material/ card weight selection (Design) | Design verification (distribution evaluation) TR-0074 | 1 | 1 | 1 | 3 | N/A | N/A | N/A | N/A | N/A |
|  |  |  |  |  |  | Geometry is design to handle anticipate load (Design) |  |  |  |  |  |  |  |  |  |  |
| 8 | Label, Inner Carton – Provides device information | Identification issues | 1 | Nuisance in preparation for the clinical application | Incomplete or ambiguous product information | Label and labeling will be designed to requirements indicated in regulation for the markets intended to be distributed (Information) | Label content review FILE-004318, FILE-008308 | 1 | 1 | 1 | 1 | N/A | N/A | N/A | N/A | N/A |
|  |  | Device use outside specification | 3 | Patient harm | Inadequate disclosure of specifications/ limitation (e.g., single-use, expiry) | Label and labeling will be designed to requirements indicated in regulation for the markets intended to be distributed (Information) | Label content review FILE-004318, FILE-008308 | 1 | 1 | 1 | 3 | N/A | N/A | N/A | N/A | N/A |
| 9 | Shipper Case – Protects packaged devices | Packaging is damaged | 3 | Device/ sterile barrier damage | Inadequate design | Material/ card weight selection (Design) | Design verification (distribution evaluation) TR-0074 | 1 | 1 | 1 | 3 | N/A | N/A | N/A | N/A | N/A |
|  |  |  |  |  |  | Geometry is design to handle anticipate load (Design) |  |  |  |  |  |  |  |  |  |  |
| 10 | Label, Shipper Case – Provides device information | Identification issues | 1 | Nuisance in preparation for the clinical application | Incomplete or ambiguous product information | Label and labeling will be designed to requirements indicated in regulation for the markets intended to be distributed (Information) | Label content review FILE-004318, FILE-008308 | 1 | 1 | 1 | 1 | N/A | N/A | N/A | N/A | N/A |

# Risk/ Benefit Assessment

## The following table indicates the risk/benefit analysis for the residual risks that were greater than the risk acceptability limit after all applicable controls/mitigations were applied:

| ### Table 7.1 – Risk/Benefit Assessment for Residual Risk above Acceptability Limit for Process Hazard Assessment |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ### Table No. | ### Process Description | ### Section No. | ### Potential Failure Mode | ### Potential Effect of Failure | ### Potential Cause of Failure | ### S | ### F2 | ### Risk/Benefit Analysis | ### Does the benefit outweigh the residual risk? |
| ### N/A | ### N/A | ### N/A | ### N/A | ### N/A | ### N/A | ### N/A | ### N/A | ### N/A | ### N/A |

## There was no residual risk identified that was above the risk acceptability limit per §5.3.

# Conclusion

## Risks have been reduced as low as possible by the following methods:

## Design the product with inherent safety

## Protective measures/ mechanisms to the product of the device.

## Provide information for safety

## Based on a review of the risk-management process for the SILQ ClearTract Foley Catheter described herein, the risk management plan based on ISO 14971:2019 has been implemented and the individual, cumulative and overall residual risk have been reduced as low as possible regardless of the magnitude and the benefit of using this device will outweigh risks posed by the device during usage.

- Change History

| ## Revision | ## Change Description |
| --- | --- |
| ## A | ## Initial release |
| ## B | ## Add HDX-160510, general updates, add verification references |
| ## C | ## Add SLQ-1610 |
| ## D | ## Update Scope Table 1, add reference to RM Plan §3.2 |